APPENDIX V

5-pH POINT TITRATION METHOD FOR THE

DETERMTNATTON OF H2CO3* ALKALINTY AND SCFA IN AQUEOUS

SOLUTIONS CONTAINING KNOWN CONCENTRATIONS OF

INORGANIC NITROGEN AND PHOSPHATE

# Introduction

The testing procedure for the 5-pH point titration method is set out for the determination of the H2CO3*alkalinity (via the total carbonate species concentration, CT, and initial sample pH) and SCFA in aqueous solutions also containing the ammonium and phosphate weak acid/bases, e.g. anaerobic digester liquid. It was developed for application in monitoring anaerobic fermentation systems. The theory of the method is set out in Chapter 5.

# Principle

The sample is titrated from its initial pH to 4 further pH points. Knowing the concentrations of inorganic nitrogen (NHs/NHi) and phosphate the concentrations of SCFA and HzCO3*alkalinity Can be derived from the theory on which the method is based. The measured data are inserted in a personal computer program to facilitate the necessary calculations (see Appendix W).

# Apparatus

The following apparatus is required (for details see Appendix B):

- titration burette (10 mL) allowing dosing increments of 0,02 mL of titrant,

- pH meter allowing readings to the second decimal p1ace,

- pH probe, preferably of the combined glass electrode type,

- a magnetic stirrer and stirrer bar (length I 25 mm),

- thermometer accurate to * 0,5 deg C,

- filter stand and ordinary filter paper, e.g. Schleicher and Schuell 505,

- measuring pipettes (A grade) from 5 to 50 mL,

- 100 mL Erlenmeyer flask

- specific conductivity meter and probe (if available), and,

- mass scale accurate to + 1 mg,

- stop watch.

# Chemicals

Hydrochloric acid, sodium hydroxide, distilled water, anhydrous sodium carbonate and pH buffer solutions (pH - 4,00 and pH: 7,00).

The hydrochloric acid requires accurate standardisation, see section "Standardization of strong acids. With the sample volumes (for anaerobic digester samples), suggested in the testing procedure below, standardisation of the strong acid to about 0,08 N is recommended, using high purity sodium carbonate. The strong base solution (if required), e.g. sodium hydroxide, should be made up to approximately the same normality as the hydrochloric acid e.g: 0,08 to 0,1 N; standardisation of the strong base is not necessary.

The pH probe needs to be calibrated using two pH buffer solutions bracketing the pH titration range of the 5-pH point titration (excluding the initial pH). It is recommended to use the following two NBS standard buffer solutions, (1) 0,05 M

potassium hydrogen phthalate buffer (pH - 4,00 at 25" C), and (2) 0,0275 M disodium hydrogen phosphate and 0,025 M potassium dihydrogen phosphate (pH 7,00 at 25" C).

# Sample preparation

A choice must be made regarding:

- Sample size (mL) undiluted.

- Sample size (mL) diluted.

Sample size (undiluted): The undiluted sample must be representative of the liquid to be tested. With an anaerobic reactor liquid usually a sufficiently large liquid volume is available for testing purpose. Filter the sample prior to the titration to (1) separate the solid from the liquid phase and (2) expel CO2 from the sample (this has the important effect in that it raises the pH of the undiluted sample, a matter of importance if the pH of the undiluted sample is below 6,7). From the filtered undiluted sample, a volume needs to be selected for titration.

The choice of undiluted sample volume is determined by (1) the dilution ratio required and (2) the diluted sample volume [see section 'sample size (diluted)" below]. Dilution of the sample is necessary for two reasons: (1) to achieve a total carbonate species concentration, C1, below about 500 mg/L as CaCO3 to avoid excessive CO2 loss during titration, (2) to reduce the total dissolved solids concentration to < 2500 mg/L to allow calculation of the activity coefficients according to the Davies Equation (see Appendix A) and, (3) to reduce the temperature of the diluted sample to about 20 - 25'C. For anaerobic reactor liquids, generally the above conditions will be satisfied using a dilution ratio of 1:4. Normally a diluted sample volume of 50 mL is adequate (see below); this implies an undiluted sample volume of 10 mL, the dilution ratio is kept low to

avoid multiplication effects of titration errors' Cr, H2CO3*alkalinity and SCFA are determined for the diluted sample by titration; the respective values for the undiluted sample are obtained by multiplying the values of the diluted sample with the dilution ratio. Bence any titration error in the diluted sample will be multiplied by the dilution ratio leading to loss of accuracy and precision.

Sample size (diluted): The diluted sample size depends on the physical properties of the pH probe and the titration vessel. In developing the 5-pH point titration method, the pHs were measured using a combined glass electrode (Radiometer Copenhagen GK2401C). The basic physical requirement to be satisfied is that the tip and porous pin (liquid junction) are immersed in the sample below the liquid surface. With the 100 mL Erlenmeyer flask as titration vessel, with the Radiometer probe, the tip and porous pin were well immersed using a diluted sample volume of 50 m/. With titration vessels greater than 100 m/, greater diluted sample volumes are required. Using Erlenmeyer flasks as titration vessels, in general, it should be noted that the surface (open to atmosphere) to volume ratio of the diluted sample should be small to minimize CO2 loss; stirring time should be short and the stirring rate gentle to minimize CO2 loss, yet adequate to achieve homogeneous mixing conditions. From practical experience these conditions appear to be satisfied with 50 mL of. diluted sample in a 100 mL Erlenmeyer flask using a magnetic stirrer bar of 25 mm length rotating at approximately 60 rpm.

# 5-pH point titration procedure

The 5-pH point titration involves the following step by step procedure:

- Pipette 10 mL of the filtered sample into a 100 mL Erlenmeyer flask and pipette 40 mL distilled water to give a dilution of 1:4. The requirement is to have C1 diluted to smaller than about 500 mg/L as CaCO3. (With the dilution ratio of 1:4 this condition usually will be satisfied for anaerobic digestor liquids. If subsequently from the analysis it is found that C1 > 500 mg/L as CaCO3 then the test needs to be repeated at a higher dilution ratio estimated from the initial test results).

- Insert thermometer and pH probe, stir gently for 15 seconds (s), take temperature reading, remove thermometer and stop stirring.

- Wait a further 45 s and record initial pH reading (this gives the probe a total stabilization period of 60 s at the initial pH).

- Switch on stirrer and titrate sample from initial pH to pH1 (6.7 + 0.1).

- When pH1 has been reached, stir for about 30 s, then switch off stirrer, record volume of titrant added from initial pH to pH1 and take pH1 reading about 30 s after termination of stirring. The volume of titrant added to titrate from pH initial to pE1 is designated V*r.

- Repeat step (4) to titrate from pH1 to pH2 (5,9 + 0,1), from pH2 to pH3 (5,2 + 0,1) and from pHs to pHa (4,3 + 0,1). The volumes of titrant recorded at each pH point (V*2, V*s, v*n) are the cumulative volumes i.e. the volume added from the initial pH point to reach the respective lower pH points.

If the initial pH of the diluted sample is below 6.7 the initial pH is recorded; strong base is added to raise the pH to pH1 6.7 + 0.1. The volume of strong base added is ignored (V*t at pE1 is set to zero when entered in the computer program). Acid titration commences from pH1, i.e. the cumulative volumes for the strong acid titration from zero to Vxz, Vxs and V*l are the strong acid volumes added to reach pHz, pHs and pHa from the starting pH point, pH1 of the initial pH of the diluted sample is 6.7 + 0,1 the initial pH equals pHi and Vxl : 0 (V*, is set to zero when entered into the computer program). The cumulative volumes for the strong acid titration V*r, V*, and V*a are the cumulative strong acid volumes added to reach pH2, pHs and pHa from the starting pH point, pH1.

Depending on the initial pH of the diluted sample, two different types of titration data tables will be obtained which can be summarized as follows:

(1) Data table for titration with strong acid only (if initial diluted sample pH > 6,7):

Vx	:	pHx

0	:	initial diluted sample pH (pHo) > 6.7

Vx1	:	pH1 (6.70 + 0.1)

Vxs	:	pH2 (5.90 + 0.1)

Vx3	:	pH3 (5.20 + 0.1)

Vx4	:	pH4 (4.30 + 0.1)

(2) Data table for titration with strong acid after addition of a strong base to raise the initial pH to pH1 or if initial pH = pH1 = 6.7 + 0.1:

Vx	:	pHx

0	:	initial diluted sample pH (pHo) = 6.7 ± 0.1

Vx1 = 0	:	pH1 (6.70 + 0.1)

Vx2	:	pH2 (5.90 + 0.1)

Vx3	:	pH3 (5.20 + 0.1)

Vx4	:	pH4 (4.30 + 0.1)

The pH values located within the + 0,1 limit ideally shor:ld be aimed for (and are easily attainable even with little experience). In some cases the pH values pH1, PHz, PH¡ and pHa might be outside the above mentioned tolerance of + 0,1 pH units, say due to rrover enthusiastic" addition of titrant. However, from experience if the sample is being "over titratedrt accidentally it is not necessary to repeat the titration provided the pE value is located within 0,2 pH udts from its ideal value.

If the sample is being I'under titratedtt, step (4) of the step by step procedure may be repeated to obtain pH1, pH2, pH3 or pH4 within the acceptable limit of + 0.1 pH units from its ideal pH value.

# Ancillary measurements

Inorganic nitrogen total species concentration to enhance the accuracy of the HzCO3*alkalinity and SCFA estimates the influence of the ammonium weak acid/base subsystem on the 5-pH point titration can be minimized by taking into account its total species concentration, N1, in the algorithm employed to calculate the HzCO3*alkalinity and SCFA. The inorganic nitrogen total species concentration can be measured according to Standard Methods (1989), (for reference see Chapter 9). If no measurement is available an approximate estimate of NT, or NT = zero is entered into the computer program. An assessment of error in measurement or neglect of this subsystem on the estimates of HzCO3* alkalinity and SCFA, is presented in Chapter 4.

Inorganic phosphate, total species concentration

Analogous to the ammonium weak acid/base subsystem the accuracy of the HzCO3* alkalinity and SCFA estimates can be enhanced by considering the total species concentration of the phosphate weak acid/base subsystem, Pr, which can be measured according to Standard Methods (1989). If no measurement is available an approximate estimate of P1, or, P1: zero is entered into the computer program. An assessment of error in measurement or neglect of this subsystem on the estimates of HzCO3* alkalinity and SCFA, is also presented in Chapter 4.

Total dissolved solids, TDS, or specific conductivity, SC, of undiluted sample

To calculate the dissociation constants of the different weak acid/bases (carbonate, SCFA, free and saline ammonia and inorganic phosphate) in the computer program, the total dissolved solids, TDS, or alternatively the specific conductivity, SC, of the undiluted sample must be entered and hence needs to be measured (see Appendix A). The measurement of one of these parameters means additional experimental effort which may not be justified seeing that the algorithm employed to determine the SCFA and HzCO3* alkalinity is not very sensitive to variations in TDS or SC. Hence approximation of these latter two parameters may be justified in many cases. Accordingly, measurement of TDS or SC may not be necessary for each filtered sample but done only daily or weekly, depending on the expected variations of the operating conditions of the reactor. This decision must be taken against the background of the accuracy desired for the determination of SCFA and H2CO3*alkalinity.

To illustrate the effect of errors in TDS on the calculation of H2CO3* alkalinity and SCFA, consider the following example: Take a typical set of titration data for an anaerobic reactor liquid sample: V, (diluted) = 50 mL, v, (undiluted) = 10 mL SCFA = 300 mg/l as HAc, H2CO3* alkalinity = 1740 mg/l as CaCO3, TDS = 330 (mg/L 21 °C (after dilution) and NT = PT = 0 mg/l; Calculate SCFA and HzCO3* alkalinity values using various hypothetical values for TDS ranging from 1000 to 7000 mg/L (in the undiluted sample). The results for these calculations are shown plotted in Fig V.1. The plot indicates that SCFA and H2CO3* alkalinity determined with the 5-pH point titration method are not sensitive to changes in TDS (or, equivalently SC).

Figure V1: Influence of changes in the total dissolved solids concentrations (TDS) on th edetermination of SCFA and H2CO3* alkalinity values for a sample containing 300mg/L of HAc and 1740 mg/L as caCO3 of H2CO3* alkalinity are calculated with hypothetical TDS concentrations ranging from 1000 to 7000 mg/L

To obtain accurate results from the 5-pH point titration method it is of great importance to determine the normality (concentration of It* ions) of the strong acid titrant ¿ccu¡ate1y. The standardization method (used in developing the 5-pH point method) is based on the theory of the H2CO3* alkalinity in an aqueous carbonate solution. The principle of the method is to titrate a solution of known H2CO3* alkalinity to the H2CO3* alkalinity equivalence point (i.e. the point of zero H2CO3* alkalinity using the strong acid to be standardised. From the mass of strong acid required, to titrate to the H2CO3* equivalence point, the normality of the strong acid is determined (see below). Using this method, one encounters the problem of identifying the H2CO3* alkalinity equivalence point.

In Chapter 2 the difficulties in identifying the H2CO3* equivalence point were highlighted and it was suggested to use the method developed by Gran (1952), (for reference see Chapter 3) to overcome the problems associated with this endpoint. The theory of Gran’s method, more specifically the First Gran Function, was discussed in detail by Loewenthal et al. (1989), (for reference see Chapter 3) and will not be dealt with here; we will deal only with the practical application of the First Gran Function i.e. calculation of the volume of strong acid added to titrate to the H2CO3* equivalence point.

The standardization of the strong acid involves the following steps: (1) preparation of strong acid, (2) Gran titration of carbonate solution, (3) application of the First Gran Function, (4) calculation of normality of strong acid.

Preparation of strong acid: In preparing the strong acid the following two aspects must be kept in mind (1) volume of undiluted sample and (2) the alkalinity of the undiluted sample. Both these factors influence the mass of strong acid required to titrate the sample from the initial pH to pH4 (in the 5-pH point titration): If the normality is high small quantities of strong acid addition will cause great changes in pH and random errors in strong acid addition will become increasingly significant. If the normality is too low large quantities of strong acid are required to effect pH changes making the titration tedious and mixing of the sample more uncertain. For the purpose of titrating a volume of 10 mL of an undiluted sample of a typical anaerobic digester a 0.08 N strong acid is recommended; to approximate this value one may dilute 35 mL of 33 percent hydrochloric acid with 5 L of CO2 free distilled water. This strong acid solution needs to be standardized using the Gran method.

Gran titration of carbonate solution: To standardise the above strong acid, a carbonate solution of known H2CO3* alkalinity is made up from CO2 free distilled water and anhydrous Na2CO3: add 5.300 g Na2CO3 to 5 L of CO2 free distilled water; take 50 mL of this and perform a Gran titration, as follows: Titrate the sample with the made up strong acid to a pH of about 3.9 and record pH value and the volume of titrant added. Titrate the sample to the final pH of about 8.4 in 4 to 5 titration steps, recording the cumulative volumes of titrant and. pH after each step to give a table of titration data, e.g.

Vx	:	pHx

13.32	:	3.90

13.34	:	3.80

13.37	:	3.70

13.40	:	3.60

13.44	:	3.50

Applying the First Gran Function: To this set of data allows the calculation of the volume of strong acid required to titrate to the H2CO3* equivalence point. Application of First Gran Function: Calculate the First Gran Function value,

Fx1	=	10 -PHx (Vs + Vx)	(v.1)

Where: 	Fx1	=	First Gran Function value,

pHx	=	pH reading in pH region 3.9 to 3.4

Vx	=	volume of strong acid required (mL) to titrate to pHx,

Vs	=	initial sample volume (mL).

This gives the following Fx1, Vx pairs:

Vx	:	pHx

13.32	:	0.0080

13.34	:	0.0100

13.37	:	0.0126

13.40	:	0.0159

13.44	:	0.0201

Plot Fx1 versus Vx; a straight line should be obtained. Draw the best straight line through the data. The interception of the line with the v, axis, v*0, (i.e. Fr,: zero) gives the volume of strong acid required to titrate to the H2CO3*equivalence point, i.e. to the point where the HzCO3*alkalinity equals zero.

(Alternative to the graphical approach V*o can be determined by applying a linear regression to the data set V* vs F*1; from the linear equation Vrq is calculated for F"1 - 0. Linear regression programs are readily available on many calculators and PC software; in the example presented here this approach was selected and V¡s was found to be 13.24 mL. With the volume of strong acid known that is required to titrate the carbonate solution of known H2CO3*alkalinity, the normality of strong acid can be calculated).

Calculation of normality of strong acid: To calculate the normality of the strong acid note that the mass of H2CO3* alkalinity added to the sample equals the mass of strong acid required to titrate to the H2CO3* equivalence point:

Ca. Vxo	=	m.Vs	(v.2)

Where 	Ca	= 	normality of strong acid (mol/L),

Vxo	=	volume of strong acid required to titrate to H2CO3* equivalence point (m/),

m	=	initial H2CO3* alkalinity concentration of sample (mol/L),

Vs	=	initial sample volume (mL)

Rearranging Eq (V.2):

Ca		=	m * (Vs/Vx0)	(v.3)

For the above example of 5.300 gNa2CO3/ (5 L distilled water), initial sample size 50 mL and

Vx0 =13.24 mL, Ca can be calculated as follows:

Ca=2*mass of Na2CO3 5,3gVolume of dist. water 5 L*molecular weight of Na2CO3 106 gmol*(VsVx0)

Note that 1 mol of Na2CO3 yields 2 moles of H2CO3* alkalinity.

Input data table for computer program

The H2CO3* alkalinity and SCFA of the undiluted sample is calculated from the titration and ancillary data obtained on the diluted sample (except for TDS and SC which are measured in the undiluted sample); a personal computer program is used for the calculations.

(The program source code is listed in Appendix W.

The input data for the computer program are as follows:

0	:	initial diluted sample pH (pHo) = 6.7 ± 0.1

Vx1 = 0	:	pH1 (6.70 + 0.1)

Vx2	:	pH2 (5.90 + 0.1)

Vx3	:	pH3 (5.20 + 0.1)

Vx4	:	pH4 (4.30 + 0.1)

pH0 (initial pH of diluted sample)

pH1 (pH after addition of Vx1)

pH2 (pH after addition of Vx2)

pH3 (pH after addition of Vx3)

pH4 (pH after addition of Vx4)

Vx1 (mL)

Y*, (mL)

V*¡ (mL)

V*r (mL)

Normality of strong acid (moL)

Sample size (undiluted) (mL)

Sample size (diluted) (mL)

Temperature (deg Celsius)

TDS (mg/L)

SC (mS/m)

NT (total species con. of free and saline ammonia, mg/L as N)

PT (total species con. of inorganic phosphate, mg/L as P)

Note that either TDS or SC needs to be entered in the program: if TDS is entered SC is calculated internally and similarly, TDS is calculated if SC is entered.

Output data table of computer program

(1) H2CO3* alkalinity of undiluted sample (mg/L as CaCO3)

(2) SCFA of undiluted sample (^ell as acetic acid)

(3) Correction for systematic pH error

Notes:

- The correction for systematic pH error is subject to the condition that the carbonate weak acid/ base dominates in the sample i.e. SCFA (mg/L as acetic acid) needs to be smaller than about 0.5 * CT (mg/L as CaCO3), see chapter 5. This condition is tested internally by the program. If the condition is not met no correction for systematic pH error is made by the program and a message is displayed in the output data table that no correction has been made; the resulting loss in accuracy in the remaining output data will however remain within acceptable limits for the purpose of monitoring of anaerobic digesters. (This situation will develop only when the digester is showing signs of failure, in which event accuracy of SCFA and H2CO3* alkalinity no longer is a prime requirement).

- The maximum correction for systematic pH efforts made by the program is set to + 0.2 pH units. If the systematic pH error is greater than + 0.2 pH units this is displayed on the screen with the recommendation that the probe should be recalibrated. From practical experience (with samples from anaerobic digesters) an average systematic pH error of about 0.05-pH units was encountered. It would appear therefore that with systematic pH errors of + 0.2 indicated by the computer program, the pH probe most likely is in error and needs to be re-calibrated, and the test repeated.

- The H2CO3* alkalinity and SCFA are the concentrations present in the undiluted sample and these are also the concentrations in the reactor. The concentration of C1 in the undiluted sample, however, is not equal to the concentration in the reactor due to CO2 loss between sampling and testing. If CT in the reactor is to be obtained, this is found from the reactor in situ pH, temperature, TDS (or SC) of the reactor liquid and the undiluted H2CO3* alkalinity determined with the 5-pH point titration method, see Chapter 2.

Example

The following example demonstrates the data input and data output (after calculation of H2CO3* alkalinity, SCFA and the systematic pH error with the aid of the computer program) for data measured on a UASB reactor effluent treating lauter tun (brewery) waste.

| Data input screen | Data |
| --- | --- |
| pH0 (initial pH of diluted sample) ..... | 8.77 |
|  | 6.80 |
|  | 5.92 |
|  | 5.35 |
|  | 4.12 |
|  | 1.40 |
|  | 3.60 |
|  | 4.56 |
|  | 4.96 |
|  | 0.0728 |
|  | 10 |
|  | 50 |
|  | 24 |
|  | 380 |

Data input screen:

pHr (pH after addition of V*J.....

pHr (pH after addition of V*r).....

pHs (pH after addition of V*s).....

pH¿ (pH after addition of V*¿) .....

V*, (mL)....

Y*, (^t) .....

V*e (m4......

v,¿ (-4.....

Normality of strong acid (mol/L).....

Sample size (undiluted) (*4.....

Sample size (diluted) (nQ.....

Temperature (deg Celsius).....

rDS (ms/m).....

sC (ms/m).....

N¡ (total species con. of free and saline ammonia, mg/L as N).

P1 (total species con. of inorganic phosphate, mg/L as P).....

8,77

6,80

5,92

5,25

4,L2

1,40

3,60

4,56

4,96

0,0728

10

50

24

380

30

25

Data output screen:

H2CO3*alkalinity of undiluted sample

@el¿as CaCO3) ......1741

SCFA of undiluted sample

@el¿ as acetic acid) ..................6

systematic pH error..............- 0,02